## Lambda-UP
Failure Rate Model | RIAC-HDBK-217Plus | Justification
Component | Intel i5-10800 | https://ark.intel.com/content/www/us/en/ark/products/199311/intel-core-i510600k-processor-12m-cache-up-to-4-80-ghz.html
Y | 2020
Beta | 0.479 | RIAC-HDBK-217Plus Table 2.2.4-1
Pi-G | 2.417e-06 | RIAC-HDBK-217Plus Section 2.2.4
Lambda-OB | 8e-06 | RIAC-HDBK-217Plus Table 2.2.4-1
DC | 0.41667 | Hypothesis that a PC is used 10 hours per day.
DC-1op | 0.28 | RIAC-HDBK-217Plus Table 2.2.4-1
Pi-DCO | 1.4881 | RIAC-HDBK-217Plus Section 2.2.4
Ea-op | 0.8 | RIAC-HDBK-217Plus Table 2.2.4-1
T-AO | 70 | Assumed 70°C as the operational temperature of the processor.
T-R | 20 | RIAC-HDBK-217Plus Table 2.2.4-1 (TR-Default)
Pi-TO | 264.6978 | RIAC-HDBK-217Plus Section 2.2.4
Lambda-EB | 0.000634 | RIAC-HDBK-217Plus Table 2.2.4-1
DC-1nonop | 0.72 | RIAC-HDBK-217Plus Table 2.2.4-1
Pi-DCN | 0.81019 | RIAC-HDBK-217Plus Section 2.2.4
RH | 0.7 | Assumed the hypothesis of an average of 70% humidity during the operation.
Ea-nonop | 0.3 | RIAC-HDBK-217Plus Table 2.2.4-1
T-AE | 20 | Assumed 20°C as the non-operational temperature of the processor (average environmental temperature).
Pi-RHT | 2.248 | RIAC-HDBK-217Plus Section 2.2.4
Lambda-TCB | 2.5e-05 | RIAC-HDBK-217Plus Table 2.2.4-1
CR | 730 | Assumed that the processor is turned on twice per day.
CR1 | 482.46 | RIAC-HDBK-217Plus Table 2.2.4-1
Pi-CR | 1.5131 | RIAC-HDBK-217Plus Section 2.2.4
DT1 | 26.5 | RIAC-HDBK-217Plus Table 2.2.4-1
Pi-DT | 48.6865 | RIAC-HDBK-217Plus Section 2.2.4
Lambda-SJB | 0.00485 | RIAC-HDBK-217Plus Table 2.2.4-1
Pi-SJDT | 2.8557 | RIAC-HDBK-217Plus Section 2.2.4
Lambda-IND | 0.000552 | RIAC-HDBK-217Plus Table 2.2.4-1
Lambda-P (FIT) | 0.014402 | RIAC-HDBK-217Plus Section 2.2.4
Lambda-P (1/h) | 1.4402e-08 | Lambda-P (FIT) / 1000000
Failure Mode | Probability | Failure Rate - Failure Mode
Perda de alimentação | 0.12 | 1.7283e-09
Saídas em aberto | 0.36 | 5.1848e-09
Saídas fixadas em nível ‘0’ | 0.09 | 1.2962e-09
Saídas fixadas em nível ‘1’ | 0.08 | 1.1522e-09
Entradas em aberto | 0.36 | 5.1848e-09

## Lambda-RAM
Failure Rate Model | RIAC-HDBK-217Plus | Justification
Component | General RAM | General-Purpose RAM
Y | 2020
Beta | 0.479 | RIAC-HDBK-217Plus Table 2.2.4-1
Pi-G | 2.417e-06 | RIAC-HDBK-217Plus Section 2.2.4
Lambda-OB | 8e-06 | RIAC-HDBK-217Plus Table 2.2.4-1
DC | 0.41667 | Hypothesis that a PC is used 10 hours per day.
DC-1op | 0.28 | RIAC-HDBK-217Plus Table 2.2.4-1
Pi-DCO | 1.4881 | RIAC-HDBK-217Plus Section 2.2.4
Ea-op | 0.8 | RIAC-HDBK-217Plus Table 2.2.4-1
T-AO | 60 | Assumed 60°C as the operational temperature of the RAM.
T-R | 20 | RIAC-HDBK-217Plus Table 2.2.4-1 (TR-Default)
Pi-TO | 128.2607 | RIAC-HDBK-217Plus Section 2.2.4
Lambda-EB | 0.000634 | RIAC-HDBK-217Plus Table 2.2.4-1
DC-1nonop | 0.72 | RIAC-HDBK-217Plus Table 2.2.4-1
Pi-DCN | 0.81019 | RIAC-HDBK-217Plus Section 2.2.4
RH | 0.7 | Assumed the hypothesis of an average of 70% humidity during the operation.
Ea-nonop | 0.3 | RIAC-HDBK-217Plus Table 2.2.4-1
T-AE | 20 | Assumed 20°C as the non-operational temperature of the RAM (average environmental temperature).
Pi-RHT | 2.248 | RIAC-HDBK-217Plus Section 2.2.4
Lambda-TCB | 2.5e-05 | RIAC-HDBK-217Plus Table 2.2.4-1
CR | 730 | Assumed that the RAM is turned on twice per day.
CR1 | 482.46 | RIAC-HDBK-217Plus Table 2.2.4-1
Pi-CR | 1.5131 | RIAC-HDBK-217Plus Section 2.2.4
DT1 | 26.5 | RIAC-HDBK-217Plus Table 2.2.4-1
Pi-DT | 26.2798 | RIAC-HDBK-217Plus Section 2.2.4
Lambda-SJB | 0.00485 | RIAC-HDBK-217Plus Table 2.2.4-1
Pi-SJDT | 2.0157 | RIAC-HDBK-217Plus Section 2.2.4
Lambda-IND | 0.000552 | RIAC-HDBK-217Plus Table 2.2.4-1
Lambda-P (FIT) | 0.010328 | RIAC-HDBK-217Plus Section 2.2.4
Lambda-P (1/h) | 1.0328e-08 | Lambda-P (FIT) / 1000000
Failure Mode | Probability | Failure Rate - Failure Mode
Perda de integridade de dados | 0.34 | 3.5115e-09
Lentidão de acesso | 0.17 | 1.7558e-09
Abertura de pinos | 0.23 | 2.3754e-09
Curto-circuito entre pinos | 0.26 | 2.6853e-09

## Lambda-CC
Failure Rate Model | RIAC-HDBK-217Plus | Justification
Component | Edge Card Connector (2 un.) | 2 connectors for RAM units
Lambda-P (FIT) | 0.00306 | RIAC-HDBK-217Plus Table 2.3-2 (multiplied by two)
Lambda-P (1/h) | 3.06e-09 | Lambda-P (FIT) / 1000000
Failure Rate Model | RIAC-HDBK-217Plus | Justification
Component | PCB Connector | Socket for Processor
Lambda-P (FIT) | 0.000168 | RIAC-HDBK-217Plus Table 2.3-2
Lambda-P (1/h) | 1.68e-10 | Lambda-P (FIT) / 1000000
Failure Rate Model | RIAC-HDBK-217Plus | Justification
Component | Rectangular Connector | Connection with Hard Drive
Lambda-P (FIT) | 0.002047 | RIAC-HDBK-217Plus Table 2.3-2
Lambda-P (1/h) | 2.047e-09 | Lambda-P (FIT) / 1000000
Lambda Total (1/h) | 5.275e-09 | Sum of Lambda-P for all connectors.
Failure Mode | Probability | Failure Rate - Failure Mode
Abertura ou aumento de resistência | 0.61 | 3.2177e-09
Curto-circuito ou diminuição de resistência | 0.16 | 8.44e-10
Perda intermitente de contato | 0.23 | 1.2132e-09

## Lambda-Unsafe
Component | Failure Mode Lambda | Safe/Unsafe | Detectable/Undetectable
CPU | 1.7283e-09 | Safe | Detectable
CPU | 5.1849e-09 | Unsafe | Undetectable
CPU | 1.2962e-09 | Unsafe | Undetectable
CPU | 1.1522e-09 | Unsafe | Undetectable
CPU | 5.1849e-09 | Unsafe | Undetectable
RAM | 3.5115e-09 | Safe | Detectable
RAM | 1.7558e-09 | Unsafe | Undetectable
RAM | 2.3754e-09 | Unsafe | Undetectable
RAM | 2.6853e-09 | Unsafe | Undetectable
PDSU | 6.6667e-07 | Unsafe | Detectable (c=0.9)
CC | 3.2178e-09 | Unsafe | Undetectable
CC | 8.44e-10 | Unsafe | Undetectable
CC | 1.2133e-09 | Unsafe | Undetectable
| 5.2398e-09 | Safe | -
| 9.1577e-08 | Unsafe | Undetectable
| 6e-07 | Unsafe | Detectable (c = 0.9)